## Slide 1
Mark H. Crosthwaite, M.Ed., CNMT, PET, NMTCB(RS), FSNMMI-TS
Associate Professor and Program Director
Virginia Commonwealth University
Generators in Nuclear Medicine

## Slide 2: Let’s Take a Look at Generators

## Slide 3: Types of Medical Reactors
Two types of fission targets are in use today: highly enriched uranium (HEU) typically containing more than 90 wt% of 235Uand low-enriched uranium (LEU) with less than 20 wt% of 235U. However, an international effort is currently underway to reduce and eventually eliminate the use of HEU targets given that they contain weapon-grade uranium [1]

## Slide 4: Methods to Produce 99Mo
https://www.ncbi.nlm.nih.gov/books/NBK215146/
100Mo(p,2n)99mTc
photo-fission of 238U
100Mo(p,2n)99M
Neutron Activation

## Slide 5: Generators and Methods of Extraction
Column chromatography – most common
  Separation technique
  Mobile phase (liquid or gas) enters the column
  Mixes with a stationary phase that has a substance(s) “fixed” to it
  As the mobile phase moves through the stationary substance separation occurs
  Required – right solvent to remove the element of interest
https://microbenotes.com/column-chromatography/
Images were modified

## Slide 6: Radiopharmaceutical Column Chromatography
Radioactive Parent (AP) is bound onto a porous
stationary phase - anionic attraction
As the AP decays a daughter is produced A-1D
  The loss of an electron (e x – 1) weakens the
  bond with the stationary phase
  A-1D – has a cationic or neutral charge
Mobile phase moves through the column
liberating A-1D
A-1D is collected in a elution vial
Ready to compound

## Slide 7: Concept --- Generators
Long-lived parent and a short-lived daughter
  Secular vs transient
The parent decays to the daughter
  Daughter builds up as parent decays
  Eluting the generator liberates the daughter
  Daughter then builds up (again) via parent decay
  Equilibrium occurs when daughter activity = parent activity
Key
  Chemical properties must be different
  Only the daughter should be in the elution
  Shielding
  Sterile
https://www.radiologycafe.com/radiology-trainees/frcr-physics-notes/production-of-radioisotopes
Modified from the original image

## Slide 8: 99Mo/99mTc Generator – Wet Column
Dry Column
Usually produced by fission of 236U
99Mo is bonded to the Al2O3.
column
99mMo decays to 99mTc it lacks affinity of the Al2O3
Usually yields >80% 99mTc yield
99Mo breakthrough should be <5.6kBq of Molly per 37 MBq of 99mTc
When saline washes across the molly/alumina core 99mTc is liberated and collected

## Slide 9: What Might Cause 99MoBreakthrough
Testing for radionuclide purity
Shield blocks low energy gammas, but not in the 740/780 keV
Reasons for Breakthrough
Too much 99Mo and not enough Al2O3 on the column
Saline pH is greater than 7
Fine particles of Al in the saline elution
Poor manufacture of the column
Excessive elations to the point that 99Mo is at the bottom of the column
I heard this happened in Chicago years ago where the column was put upside down

## Slide 10: Chemical Purity – Colorimetric Test
Alumina breakthrough is tested by applying a standard that contains 10µg/mL
This litmus test changes white if the elution drop is 10µg/mL or >
High levels of alumina can effect the tagging of 99mTc to sulfur colloid and labeled red cells
Alumina can also cause flocculation – note RES uptake from radio-tagged Alumina
Question – If 99Mo breaks though would the alumina?

## Slide 11: RadioGenixTM Liquid 99Mo Generator
https://physicstoday.scitation.org/do/10.1063/PT.6.2.20180216a/full/
99Mo is made by neutron activation hence the mixture also contains 98/100 Mo all in liquid form
99Mo is pumped into PSC cabinet where it passes though the chromographic column but 99mTc adheres to the column
99Mo eventually returns to the source cabinet
99mTc remains on the first column then washed with specific solutions and then sent to the Production Cabinet.
0.9% saline removes 99mTc from the column and it is ready for compounding
Waste is then sent to the Discard cabinet

## Slide 12: 82Sr/82Rb Generator System
85Rb(p,4n)82Sr
T1/2 82Sr = 25 days, 85Sr = 64.8 days, and 82Rb – 1.25 minutes
Typical generator has between 3.7GBq to 5.55 GBq of 82Sr activity
Stationary phase is made of SnO2 that binds 82Sr and 85Sr to the column
82Rb yield is 40 to 90% with a pH of 7
pH factor – lower pH (<7) releases greater amounts of Sr while higher levels (9) reduce Sr, but increase Rb
Higher infusion rates increase the % yield of 82Rb and 82/85Sr
Generator equilibrium takes <10 minutes

## Slide 13: 82Sr/82Rb Generator System

## Slide 14: 82Sr Breakthrough
From generator elution calculate back
to time zero for 82Rb
Determine R
  Generator label has the 85/82Sr ratio
  T1/2 82Sr = 25 days 85Sr = 64.8 days
  Define ratio factor based on decay
Now determine 82Sr level
  Dose calibrator reading
  Plug in R
  F is always 0.478
  Calculate
Now you’re ready to apply

## Slide 15: Too Much 82/85SrBreakthrough
Summer 2011 two individuals triggered sensors at the Canadian Board – 82Sr and 85Sr
Generators called – 113 tested negative
New QC implemented
New breakthrough limits (daily)
    Less than one-tenth of 0.02µ (82Sr/85Sr) per mCi of 82Rb
    If limit is reached complete breakthrough test x2 per day
  Breakthrough to be done if 750 mL passes through the system in one day
Use additive-free Sodium Chloride
Breakthrough is >0.02µCi with 82Sr or >0.2 with 85Sr discontinue use and contact the distributor

## Slide 16: 68Ge/68Ga Generator
68Zn(p,n)68Ge
Chromatographic column surrounded by glass with a TiO2 core
  68Ge bonds to the TiO2 and has a T1/2 of 271 days
  68Ga is the daughter with a 68 minutes
Elute with 0.1M HCl in 5 mL
68Ga yield ~75% in a 3 mL elution
68Ge breakthrough is <0.005%
You are now ready to tag it to DOTATATE or PMSA
In US PMSA is in clinical trials

## Slide 17: 81Rb/81mKr Generator
84Kr(p,4n)81Rb or 85Rb(p,p4n)81Rb
81Rb is bonded to a cation-exchange resin (Dowex50wx8)
Generator activity ranges from 111MBq to 925 MBq at calibration
Elution – lungs (gas phase) or RV (liquid phase)
Moist air or O2 (at 1 L/min) liberates the 81mK or continues elution with IV administration (5% glucose)
0.9% NaCl solution yields similar extraction results
81mKr yields 90% - breakthrough “<10-8/mL”
Apparently still made by CyclotronVU
My experience – SLO after 3pm on any given day

## Slide 18: Honorable Mention– How About a Few More Generators

## Slide 20: References
Chattopadhyay, S. et al. (2019) A compact solvent extraction based 99Mo/99 99mTc generator for hospital radiopharmacy. (online) https://www.sciencedirect.com/science/article/pii/S0969804318307619 APPL RADIAT ISOTOPES 143:41-46 (accessed 5 Oct 2019)
Skliarova H. (2019) Recovery of Molybdenum Precursor Material in Cyclotron-Based Technetium-99m Production Cycle. (online) https://www.mdpi.com/2410-390X/3/1/17/htm Instruments (accessed Oct 5 2019)
Sharma S. (2014) Radionuclide Generators. (online) https://www.slideserve.com/umika/radionuclide-generators. Depart of Nuc Med, AIIMAS PPT (accessed Oct 5 2019)
Package Insert Technetium Tc99m Generator (online) https://www.accessdata.fda.gov/drugsatfda_docs/label/2013/017693s025PI.pdf NDA 17693/S-025 (accessed Oct 5 2019)
Cyclotron Produced Radionuclides: Physical Characteristics and Production Methods ( (online) https://www-pub.iaea.org/MTCD/publications/PDF/trs468_web.pdf Technical Report Series #468 2009) IAEA. (accessed 6 Oct 2019)
Mausner, LF, et al. The Production of 82Rb/81mKr Generators with =60-MeV Protons at BLIP (online) https://www.osti.gov/servlets/purl/5944453 International Symposium on Single-Photon Ultra-Short-Lived Radionuclides (accessed 6 Oct 2019)
Dulugeac S.D., (2013) Neutronic Analysis For The Fission Mo99 Production By Irradiation of LEU Targets INTRIGA 14 MW Reactor. (online) Institute of Nuclear Research, Pitesti, Romania https://inis.iaea.org/collection/NCLCollectionStore/_Public/47/053/47053341.pdf?r=1&r=1 (accessed 18 Sept 2019)
Phillips D. (1996) Radionuclide generator systems for nuclear medicine. (online) https://pharmacyce.unm.edu/nuclear_program/freelessonfiles/Vol5Lesson1.pdf Univ of NM College of Pharmacy. Volume V, Number 1. (accessed Oct 5 2019)
Fennell E.J., The RadioGenixTM System. (online) https://apha2018.pharmacist.com/sites/default/files/slides/Radioisotope_Procurement_Fennell_3-19-18_105A_HO.pdf (accessed 19 Sept 2019)
Zoller, F. (2006) The Role of Copper in Nuclear Medicine. NCSS (online) https://www.slideshare.net/brucelee55/the-role-of-copper-in-nuclear-medicine (accessed 19 Sept 2019)

## Slide 21: References
Assessment Report. (2012) Cuprymina (online) https://www.ema.europa.eu/en/documents/assessment-report/cuprymina-epar-public-assessment-report_en.pdf (assessed 3 Oct 2019)
Sylvester P, et. al. (2004) Method of separation of 90Y from 90Sr. Priority 2002 -06-18 (online) https://patentimages.storage.googleapis.com/e8/73/95/056650ffbcfb37/US20040005272A1.pdf (accessed 3 Oct 2019)
Injarean U, et al. (2018) Process for separation of Y-90 from SR-90 in HNO3using combined solvent impregnated with resins of D2EHPA/dodecane and CMPO/TBP Chiang Mai Journ of Sc 45(1):556-564 (online) https://www.researchgate.net/publication/323837109_Process_for_separation_of_Y-90_from_Sr-90_in_HNO3_using_combined_solvent_impregnated_resins_of_D2EHPA_dodecane_and_CMPOTBP (accessed 3 Oct 2019)
Haynes, NG, et. al. (2000) Performance of a 62Zn/62CuGenerator in Clinical Trials of PET Perfusion Agent 62Cu-PTSM (online) http://jnm.snmjournals.org/content/41/2/309.full.pdf+html?sid=03967db9-5f98-40ba-a4a2-15340bb37f90 J Nucl Med; 41:309-314 (accessed 4 Oct 2019)
Beyer, GJ. (1990) A 81RB/81mKr generator made by ion implantation (online) http://cds.cern.ch/record/210342/files/cer-000122319.pdf European Organizations of Nuclear Research (accessed Oct 4 2019)
Kato M, Hazue (1977) [Generator kinetics of a 81mKr-generator for medical use [author’s transl] (online) https://www.jstage.jst.go.jp/article/radioisotopes1952/26/3/26_3_179/_pdf Radioisotopes 26(3):179-181 (accessed 4 Oct 2019)
Dash A, et. Al. (2019) Radionuclide generators: the prospect of availing PET radiotracers to meet current clinical needs and future research demands (online) http://www.ajnmmi.us/files/ajnmmi0086228.pdfnAM J Nucl Med Mol Imaging 9(1):30-36 (accessed 4 Oct 2019)
Suleiman, OH (2012)FDA Update: The Strontium/Rubidium Generator (online) https://www.nrc.gov/docs/ML1210/ML12101A274.pdf Presented NRC Advisory Committee (Accessed 6 Oct 2019)
Cherundolo B. 2014 Meeting of Mid-Atlantic States Radiation Control Program. (online) https://manualzz.com/doc/6907637/bracco-s-rb-82-generator Bracco website (accessed 6 Oct 2019)